## Table of Contents
Table of Contents
Sheet Number | Link to Tab
1 | Table of Contents
2 | App Dev Guide DocVoice
3 | Error Codes
4 | BLE GATT Characteristics
5 | Device Settings Byte Structure
6 | State Table
Confidentiality Notice: This document is intended solely for the internal use of DocVoice and its parent company. It contains proprietary and confidential information that must not be shared, distributed, or disclosed to any external parties without prior written consent from Salvica.
Release Date | Oct 23rd, 2025
Document Version | 1.1.0 A
Firmware Version | 1.1.0

## App Dev Guide DocVoice
If using nRF Connect for BLE Testing, ensure all reading/writing is done in ASCII format (UTF-8) for the responses to be correctly parsed.
Mobile App Actions | Device State Response in ASCII format | LED status indicator
Scan for nearby BLE device. If user wants to connect to a new device, check the device BLE advertising packet: Company Identifier: 0xFFFF Data: _ "SAL1" if this device is publicly advertised (meaning anyone can pair) _ "SAL0" if this device is advertised to only paired device (meaning the device belongs to someone else) App should rely on the prefix "SAL" in advertising packet to filter only your device (hide other nearby BLE devices), and only show the device that is publicly advertising when scanning for new device. This will be helpful to filter out the board that is already paired to other smartphones, thus isn't in public advertise mode. | N/A | Blue LED blinks fast if publicly advertise Otherwise blink slow
Connect to BLE | "s2" | Blue LED blinks fast if publicly advertise Otherwise blink slow
Subcribe to notification of: 1. All characteristic in Battery Service and Device Information Service 2. User Control characteristic of the Custom Service 3. Vitals and Temperature Characteristics | "s2" | Blue LED blinks fast if publicly advertise Otherwise blink slow
Read the device data: 1. Battery Level (this is notified every 10s, but app should still read it as soon as connected, to show battery level immediately to user) 2. Model Number 3. Serial Number 4. Firmware Revision 5. Hardware Revision 6. Manufacture Name | "s2" | Blue LED blinks fast if publicly advertise Otherwise blink slow
In this version of the firmware, the User Control characteristic expects a One Letter Tag and a value, in this order. The One Letter Tag corresponds to an input type, the value is what that input read or write: One Letter Tags (first byte in command): 'k' - Passkey 'c' - Command 'e' - Error 't' - Settings 's' - Status
Ensure device status = "s2", write passkey "k1234" in ASCII format to User Control characteristic. This step is mandatory to proceed. | "s3" Current Device Settings | All LED are off
Device always return its current settings upon entering correct passkey for synchronization purpose. If the user settings are different from current device settings, send the up-to-date device settings from the app to the device via User Control characteristic. Only send settings while being in "s3" state. Sending setting in different states will simply be ignored. | "s3" | All LED are off
Send "c1" to User Control characteristic to start the dataset. Device status will change to "s4" to initialize all sensors, then quickly change to: 1. If state = "s5", error ocurred. Check value return from User Control characterstic to determine the type of error. Then send "c2" to User Control characteristic to get back to state "s3" before trying to start the dataset again 2. If state ="s6", device is not properly placed on the user skin. App can show a instructional dialog for repositioning the device. There is no data being collected at this point yet | "s4" then quickly changes to "s5" or "s6" | Red LED blinks fast if state = "s5"
All LED are off if state = "s6"
If user has placed the device on the skin properly, state will change from "s6" to "s7" shortly. Once in "s7" state, data will be continuously sent to the smartphone via Vitals Data and Temperature characteristics. Refer to Bluetooth GATT Version Controlon how to decode the data | "s7" | All LED are off
Once the dataset duration is elapsed, device will sleep; however, BLE connection is still sustained | "s8" | All LED are off
Once the sleep duration is elapsed, device will wakeup and automatically start dataset and the loop goes on | ("s8"->"s4"->"s6"->"s7") | All LED are off
At any time during the dataset or sleep, app can send "c2" to User Control characteristic to stop the dataset/sleep cycle | "s3" | All LED are off
At any time during the dataset, if red PPG signal strength is below off skin threshold, dataset willl be automatically terminated | "s3" | All LED are off
At any time, app can send "c3" to User Control characteristic to erase all data. This should only be used when user wants to give the device to another person. BLE pairing data is cleared. App should also clear the pairing data of the board on smartphone. | "s1" | Blue LED blink fast as it will publicly advertise for new user
At any time during connected, device will go to sleep after: _60 seconds if no correct passkey is received during "s2" state _300 seconds if dataset isn't started during "s3" state _60 seconds if skin isn't detected during "s6" state | "s0" | All LED are off
At any time, device will go to sleep immediately if any condition below is true: _ battery level reading is 0% _MCU temperature is above 50 degree Celcius |  | All LED are off
Confidentiality Notice: This document is intended solely for the internal use of DocVoice and its parent company. It contains proprietary and confidential information that must not be shared, distributed, or disclosed to any external parties without prior written consent from Salvica.

## Error Codes
Each bit represents a specific error condition. A bit is set to 1 if the corresponding error occurs; otherwise, it remains 0.
Group | Error Description | Error occurance | Reason of error | Bit position
Vitals Sensor (PPG sensor) | No response | Cannot initialize the sensor through I2C communication | Misplacement of the Vital Sensor IC | 0
Board | Supply voltage out of range | Power issues | Supply not controlled to specified range. | 1
Temperature Sensor | Temperature sensor can't be initialized | Cannot initialize the sensor through I2C communication | Misplacement of the temperature sensor | 9
Settings | Invalid Length | Received data length in Settings cahracteristic is incompatible | Wrong data format for device settings | 11
| Invalid Value | Received value of at least 1 field in Settings characteristic is out of range | Wrong data format, wrong big/little endian | 12
Command | Invalid | Command is not supported in Command characteristic | Typo in command by the mobile app dev Command is not supported in a specific state | 13
Bluetooth | Packet Drop | Bluetooth data packet queue is full | The Bluetooth signal quality is too weak to handle the required data rate App should prompt user to move closer to the smartphone for better signal quality | 14
Data | Weak PPG Signal | During dataset, PPG signal strength drops below threshold specified in proximity threshold | Too big airgap between PPG sensor and the skin of user | 15
MCU | MCU overheating | Can be triggered at anytime if MCU temperature exceeds 50 degrees C Device will turned off immediately after sending this error | Battery overheating Ambient temperature is too high | 17
Confidentiality Notice: This document is intended solely for the internal use of DocVoice and its parent company. It contains proprietary and confidential information that must not be shared, distributed, or disclosed to any external parties without prior written consent from Salvica.

## BLE GATT Characteristics
LEGEND
|  | * Write: sending data from smartphone to the board |  |  |  | NO LEDS
|  | * Notify: smartphone received data from the board via callback |  |  |  | INAPPLICABLE FUNCTIONS/PROPERTIES
|  | All notify requires BLE connected |  |  |  | PASSKEY REQUIRED
BLE Default Name: SAL-IDNUMBER | <- this name will change | Firmware Version: 1.1.0 | <- this name will change
Service | Characteristic | UUID | Properties | Tags | Description | Write | Value | Notify timing | LED status
Battery Service (0000180f-0000-1000-8000-00805f9b34fb) | Battery Level | 00002a19-0000-1000-8000-00805f9b34fb | read Notify |  | Battery level |  | Battery percentage in byte format | every 10 seconds or once enter s9 state
Device Information (0000180a-0000-1000-8000-00805f9b34fb) | Model Number String | 00002a24-0000-1000-8000-00805f9b34fb | read |  | Board model |  | "DocVoice Devboard" | once after connected
| Serial Number String | 00002a25-0000-1000-8000-00805f9b34fb | read |  | Board ID number |  | "DDMMYYIDID" | once after connected
| Firmware Revision String | 00002a26-0000-1000-8000-00805f9b34fb | read |  | Board's current firmware revision |  | "1.0.0" | once after connected & after OTA update
| Hardware Revision String | 00002a27-0000-1000-8000-00805f9b34fb | read |  | Board's hardware revision |  | "1.0" | once after connected
| Manufacture Name String | 00002a29-0000-1000-8000-00805f9b34fb | read |  | Board's manufacturer name |  | "Salvica Health" | once after connected
Custom Service (00000000-0000-0000-0000-000000000000) | User Control | 00000001-0000-0000-0000-000000000000 | Read, Write, and Notify | STATUS_TAG = 's' | Off |  | "s0" | After issuing "c0" command
[5 empty] | Disconnected |  | "s1" | app should never receive this value because it is sent after BLE is disconnected It is solely for firmware state machine | Blue LED blinks fast if publicly advertises, otherwise blink slow
[5 empty] | Unverified |  | "s2" | After BLE connected
[5 empty] | Ready |  | "s3" | correct passkey is submitted
[5 empty] | Initializing sensors |  | "s4" | After issuing "c1" command
[5 empty] | Sensor Error |  | "s5" | At least 1 error is occurred during "s4" state. Read Error Code characteristic for specific error code | Red LED blinks fast
[5 empty] | Cannot Detect User's Skin |  | "s6" | After "s5" state and sensor couldn't detect user skin | Red LED blinks slow
[5 empty] | Working/Data collection |  | "s7" | After "s6" state and device has detected user skin
[5 empty] | Sleep between dataset |  | "s8" | current dataset is finished, device will sleep until next dataset. BLE connection is sustained
|  |  |  | ERROR_TAG = 'e' | Error code such as sensor error, settings error |  | 32 bit number (little endian), referring to Error Code | _Once a new error occurred in "s4" state _While updating device settings _Invalid Command
|  |  |  | KEY_TAG = 'k' | Enter passkey upon connection | "k1234"
|  |  |  | COMMAND_TAG = 'c' | Start the dataset (can only use in "s3" state) | "c1"
[5 empty] | Stop the current dataset (can only use in any state between "s4" and "s8") | "c2"
[5 empty] | Factory reset (can use at anytime) | "c3"
[5 empty] | Turn off the device (can use at anytime) | "c0"
|  |  |  | SETTING_TAG = 't' | Changing settings of device (only when device state is Ready) | Device Settings Byte Structure (little endian)
| Vitals Data | 00000002-0000-0000-0000-000000000000 | Notify |  | PPG |  | Each sample always contains 9 bytes. Each sample is PPG (IR, Red, Green) In that Order Data is left Justified | about every 67.5 ms
| Temperature | 00000003-0000-0000-0000-000000000000 | Notify |  | Temperature data |  | Degree Celcius but being multiplied by 100 to be an integer 0x0E38 -> 3640 = 36.40C degree | start command is issued in Command characteristic, every 10 seconds. Update frequency can be changed in Settings
Confidentiality Notice: This document is intended solely for the internal use of DocVoice and its parent company. It contains proprietary and confidential information that must not be shared, distributed, or disclosed to any external parties without prior written consent from Salvica.

## Device Settings Byte Structure
Field Length (bytes)
10 + (8 to 37), min. 18 bytes total
Note: All Settings are received in Little Endian (Least Significant Byte First) for any Field that is longer than one Byte.
Factory reset values are indicated in this color
Index | Field | Field Description | Field Type | Field Byte Length | Data Value in DEC | Data Value in HEX | Data Value Description
0 | SpO2 Sample Rate Control | The effective sampling rate with one sample consisting of one IR pulse/conversion, one RED pulse/ conversion, and one GREEN pulse/conversion. The sample rate and pulse-width are related in that the sample rate sets an upper bound on the pulse-width time. If the user selects a sample rate that is too high for the selected LED_PW setting, the highest possible sample rate is programmed instead into the register. | uint8_t | 1 | 0 | 0x00 | 50Hz
[5 empty] | 4 | 0x04 | 100Hz
[5 empty] | 8 | 0x08 | 200Hz
[5 empty] | 12 | 0x0C | 400Hz
[5 empty] | 16 | 0x10 | 800Hz
[5 empty] | 20 | 0x14 | 1000Hz
[5 empty] | 24 | 0x18 | 1600Hz
[5 empty] | 28 | 0x1C | 3200Hz
1 | LED Pulse Width Control | The LED pulse width (the IR, Red, and Green have the same pulse width), and, therefore, indirectly sets the integration time of the ADC in each sample. The ADC resolution is directly related to the integration time. In micro seconds. | uint8_t | 1 | 0 | 0x00 | 69us
[5 empty] | 1 | 0x01 | 118us
[5 empty] | 2 | 0x02 | 215us
[5 empty] | 3 | 0x03 | 411us (cannot be used for 400Hz sampling rate)
2 | SpO2 ADC Range | SpO2 sensor ADC’s full-scale range | uint8_t | 1 | 0 | 0x00 | LSB Size (pA): 7.81 Full Scale (nA): 2048
[5 empty] | 32 | 0x20 | LSB Size (pA): 15.63 Full Scale (nA): 4096
[5 empty] | 64 | 0x40 | LSB Size (pA): 31.25Full Scale (nA): 8192
[5 empty] | 96 | 0x60 | LSB Size (pA): 62.5Full Scale (nA): 16384
3 | Sample Averaging | To reduce the amount of data throughput, adjacent samples (in each individual channel) can be averaged and decimated on the chip by setting this register Higher value may require depending on the LED specification, but higher power consumption. | uint8_t | 1 | 0 | 0x00 | 1 (No Averaging)
[5 empty] | 32 | 0x20 | 2 samples averaged per FIFO sample
[5 empty] | 64 | 0x40 | 4 samples averaged per FIFO sample
[5 empty] | 96 | 0x60 | 8 samples averaged per FIFO sample
[5 empty] | 128 | 0x80 | 16 samples averaged per FIFO sample
[5 empty] | 160 | 0xA0 | 32 samples averaged per FIFO sample
4 | Proximity LED Current | Proximity LED Current Control from 0mA to 51mA. | uint8_t | 1 | 0-255 | 0x00-0xFF | Default is 5 (1mA) to set LED Drive Current Range
5 | Red LED Current | Red LED Current Control from 0mA to 51mA. | uint8_t | 1 | 0-255 | 0x00-0xFF | Default is 25 (5mA) to set LED Drive Current Range
6 | IR LED Current | IR LED Current Control from 0mA to 51mA. | uint8_t | 1 | 0-255 | 0x00-0xFF | Default is 25 (5mA) to set LED Drive Current Range
7 | Green LED Current | Green LED Current Control from 0mA to 51mA. | uint8_t | 1 | 0-255 | 0x00-0xFF | Default is 150 (30mA) to set LED Drive Current Range
8 | Proximity Threshold | The proximity threshold level for the adc reading. Later it is scaled up to match the 18-bit ADC resolution by multiplying by 1024. Final threshold can range from minimum 0 to maximum 0x3FFFF. | uint8_t | 1 | 0-255 | 0x00-0xFF | Default is 10, which is 10*1024=10,240 in ADC reading
9 | Off Skin Threshold | The skin detection level for the adc reading. Later it is scaled up to match the 18-bit ADC resolution by multiplying by 1024. Final scaled threshold can range from minimum 0 to maximum 0x3FFFF. | uint8_t | 1 | 0-255 | 0x00-0xFF | Default is 50, which is 50*1024=51,200 in ADC reading
10 | Temperature Update Period | The amount of second between each temperature update of characteristic | uint16_t | 2 | 1-65536 | 0x0001-0xFFFF | Default is 0x000A or 10, which is 10 seconds between updates of temperature characteristic.
11 | Dataset Duration | The length of dataset in second | uint16_t | 2 | 0-65536 | 0x0000-0xFFFF | Default is 300, set high for long live data recording
12 | Sleep Duration | The length of sleep between dataset in second | uint16_t | 2 | 0-65536 | 0x0000-0xFFFF | Default is 300
13 | TX power | Transmit power of Bluetooth. Higher value gives stronger Bluetooth signal quality, but higher power consumption | uint8_t | 1 | 0 | 0x00 | -40dBm
[5 empty] | 1 | 0x01 | -20dBm
[5 empty] | 2 | 0x02 | -16dBm
[5 empty] | 3 | 0x03 | -12dBm
[5 empty] | 4 | 0x04 | -8dBm
[5 empty] | 5 | 0x05 | -4dBm
[5 empty] | 6 | 0x06 | 0dBm
[5 empty] | 7 | 0x07 | 3dBm
[5 empty] | 8 | 0x08 | 4dBm
14 | LED indicator brightness level | The brightness of the indicator LED (for both blue and red) | uint8_t | 1 | 0-255 | 0x00-0xFF | Default is 0xFF
15 | Device name | The device Bluetooth name | String | 29 |  |  | Each byte is each character. Max name length is 29. This field can be left empty, which means using factory default name. Set as DEVICE_SIGNATURE DEVICE_ID_DEFAULT, where DEVICE_SIGNATURE is SAL- and DEVICE_ID_DEFAULT is, for example, 2909250001
Confidentiality Notice: This document is intended solely for the internal use of DocVoice and its parent company. It contains proprietary and confidential information that must not be shared, distributed, or disclosed to any external parties without prior written consent from Salvica.

## State Table
State | Name | Description
s0 | Sleep System Off | Device is turned off
s1 | Disconnected | No BLE Connection
s2 | Waiting for correct passkey | BLE connected but no correct passkey is received
s3 | Ready | Passkey is verified and the device is ready to receive command
s4 | Initializing sensors | Initializing sensors before starting a dataset
s5 | Sensor Initialization Error | At least 1 sensor failed to initialize
s6 | No Contact of skin | The sensor isn't connected to user's skin
s7 | Working/Collecting data | Collecting and transfering vital data via BLE to smartphone or BLE hub
s8 | Sleeping between dataset | Device sleep between datasets
Confidentiality Notice: This document is intended solely for the internal use of DocVoice and its parent company. It contains proprietary and confidential information that must not be shared, distributed, or disclosed to any external parties without prior written consent from Salvica.

## Firmware Revision History
Firmware History
Version | Changes | Release Date
1.1.0 | _New BLE Structure, merge State,Command,Error,Settings into 1 User Control characteristic _Send current settings to peer device upon correct passkey verification for synchronization _Send state immediately when user listen to notification on control characteristic | 2025-10-23
1.0.0 | Initial release | 2025-10-10
Document Revisions
Revision | Changes | Release Date
1.1.0 A | Update document BLE GATT Characteristics for firmware ver 1.1.0 | 2025-10-23
1.0.0 B | Cleaned up typos | 2025-10-16
1.0.0 A | Initial release | 2025-10-10